Figure 1: Start overview

1. Select ‘Patient’ Management (Figure 1)

Figure 2: Setup Patient

2. Select ‘New’

3. Add ‘First name’ and ‘Last name’: ‘NumberParticipant’ x2 (Figure 2)

4. Add ‘Date of birth’: 1/1/1999’ for everybody (default) (Figure 2)

5. Leave ‘Gender’ as is on ‘Female’ (default) (Figure 2)

6. ‘Remark’: ADD ‘DOB and Gender are classified’ (Figure 2)

Figure 3: Group setup

7. Go to ‘Group’ – ‘New’ and add ‘Name’: ‘WP3: Healthy participants’ (Once)

8. Select participant – Go to ‘Edit’ roll down – ‘Groups’ – ‘Assign Group’ (roll down)

Figure 4: Settings

9. Go to ‘Settings’

Figure 5: Chips

10. Go to ‘chips’

Figure 6: Settings – chips

11. Settings

- Unit: ‘Celsius’

- Hours/day: ‘8’ (Desired standard wearing time)

- Minimum continuous hours: ‘0’ (Default)

- Temperature range: INSOLE (25  38.5)

Remark: can be individualised (not applicable in this clinical trial)

Figure 7: USB-BT- Pen Reader

Figure 8: How to activate the sensor

12. Press ‘ON’

13. Connect the Pen Reader to a USB port or select the Bluetooth icon (pen needs to be charged)  (Figure 7)

14. Place the tip of the Pen on the sensor: on the ‘O’ of the Orthotimer logo (Green or Orange led lights up)

15. Open the desktop app

16. Look if there is a connection

17. Touch the sensor with the pen  the red bar changes to green (Sensor ‘recognized’)

Figure 9: Activation

18. Press ‘Activate’ (Figure 9)

Figure 10: Settings Sensor

19. Select Type: ‘Insole’

20. Interval: ‘900 seconds’ (15 minutes)

Figure 11: Remark

The sensor is now activated and must perform a test measurement. For the purpose the senor should be removed from reader for a little longer than the measurement interval set.

Figure 12: Test

21. Hold the reader on the sensor and press ‘Test’

22. After the first test measurement is downloaded from the senor, a confirmation appears that the test was successful.

Figure 13: Register

23. Now the sensor will be registered and connected to the patient data. Hold the reader to the sensor and click on Register.

24. Press ‘Yes’ when assigning to participants.

25. Press the ‘*’ in the right upper corner to register a new participant. If already created, select in the search window  and Press ‘OK’ (Figure 14)

26. When the registration confirmation appears, click ‘OK’

Figure 14: Reader Connected

Here you get the connection to the Orthotimer Cloud

Figure 15: Chipmanagement

27. Go to ‘Chipmanagement’

28. Select the participant

Figure 16: Details

29. Go to roll window next to ‘Details’ and select ‘Export raw data’.